Hardware Architecture & Schematic Design

Low-Cost Smart Pen System for Parkinson's Disease Tremor Screening

ESP32-S3 SuperMini  ·  MPU-6050  ·  Tri-FSR Array  ·  Option 2 Boot-Safe RC Decoupling

Version 2.0  ·  June 2026

| ! | DISCLAIMER: This system is a non-diagnostic screening and monitoring tool. It is not cleared as a medical device. All output must be interpreted alongside professional clinical assessment using validated scales (e.g. MDS-UPDRS). |
| --- | --- |

# 1. System Overview

The hardware architecture balances high-fidelity multi-channel data acquisition, edge-computing performance, and strict spatial constraints imposed by the pen form factor. The system comprises three primary sub-circuits: the Central Processing Subsystem (ESP32-S3 SuperMini), the Sensor Acquisition Network (MPU-6050 + Tri-FSR array), and the Power Management Subsystem (LiPo + LDO). All subsystems share a single stable 3.3 V regulated rail to prevent analog FSR readings from being corrupted by battery discharge variation.

## 1.1 Signal Pipeline

Sensor Acquisition  →  ADC1 sampling @ 1 kHz  →  DC offset removal

→  High-pass filter (1 Hz)  →  Band-pass filter (3.5–6.5 Hz)

→  1024-point FFT (on-device)  →  Feature extraction

→  Native BLE transmission @ 1 Hz  →  Mobile application

## 1.2 Component Summary

| Subsystem | Component | Key Details |
| --- | --- | --- |
| Central processing | ESP32-S3 SuperMini | Dual-core LX7 @ 240 MHz · native BLE 5 · ADC1 channels · onboard LDO |
| Kinematic sensing | MPU-6050 MEMS IMU | I2C digital · 4–6 Hz tremor range · 100 nF VCC decoupling · AD0 → GND (0x68) |
| Axial force sensing | FSR-400 (tip) | 7.5 mm active area · plunger-isolated · Option 2 RC circuit on GPIO 1 |
| Grip force sensing A | FSR-402 (barrel A) | 44 mm active area · fingerprint zone · Option 2 RC circuit on GPIO 2 |
| Grip force sensing B | FSR-402 (barrel B) | 44 mm active area · fingerprint zone · Option 2 RC circuit on GPIO 3 |
| Power | LiPo 3.7 V · 500 mAh | 4–6 hr runtime · onboard LDO → flat 3.3 V rail for ADC stability |

# 2. Central Processing Subsystem — ESP32-S3 SuperMini

The ESP32-S3 SuperMini is an ultra-compact development board housing the dual-core Xtensa 32-bit LX7 microprocessor. It was selected over standard ARM Cortex-M or 8-bit microcontrollers for three reasons: sufficient computational headroom for 1 kHz multi-channel sampling with concurrent FFT processing, native BLE 5 eliminating the need for an external UART module, and an onboard LDO providing the stable 3.3 V rail essential for reliable analog FSR readings.

## 2.1 Key Specifications

| Parameter | Value |
| --- | --- |
| MCU | ESP32-S3 (Xtensa dual-core LX7) |
| Clock speed | Up to 240 MHz (dual-core) |
| Operating voltage | 3.3 V logic (regulated by onboard LDO) |
| Wireless | Native BLE 5 · 2.4 GHz onboard PCB antenna · no external module required |
| ADC resolution | 12-bit nominal · ~11-bit effective with noise |
| ADC channels used | ADC1 only (GPIO 1, 2, 3, 4) · ADC2 avoided — incompatible with active BLE |
| I2C | GPIO 8 (SDA) · GPIO 9 (SCL) |
| Flash / PSRAM | 8 MB flash · 2 MB PSRAM (model-dependent) |

## 2.2 GPIO Pin Assignment

Pins are partitioned to separate the high-speed I2C digital bus from the sensitive analog ADC channels, preventing crosstalk-induced noise on FSR traces. ADC2 channels are avoided entirely as they share internal hardware paths with the BLE radio stack and produce unreliable readings during active wireless transmission.

| GPIO | Function | Interface | Notes |
| --- | --- | --- | --- |
| GPIO 8 | I2C SDA | MPU-6050 data | 4.7 kΩ pull-up to 3.3 V required |
| GPIO 9 | I2C SCL | MPU-6050 clock | 4.7 kΩ pull-up to 3.3 V required |
| GPIO 1 | ADC1_CH1 | FSR-400 tip sensor | Strapping pin — Option 2 RC mitigation applied |
| GPIO 2 | ADC1_CH2 | FSR-402 grip A | Strapping pin — Option 2 RC mitigation applied |
| GPIO 3 | ADC1_CH3 | FSR-402 grip B | Strapping pin — Option 2 RC mitigation applied |

| ! | STRAPPING PIN WARNING: GPIO 0, 1, 2, and 3 are ESP32-S3 strapping pins sampled during boot to determine startup mode. A fixed voltage divider connected directly to these pins can hold an incorrect voltage level during boot, causing boot loops or preventing firmware upload. Option 2 RC decoupling (Section 4) is mandatory on all three FSR channels to mitigate this risk. |
| --- | --- |

# 3. Sensor Acquisition Network

The sensor network uses two distinct electrical interfaces: a digital I2C bus for the MPU-6050 IMU and a passive analog voltage divider front-end for each of the three FSRs. The two interface types are electrically isolated through GPIO partitioning to prevent digital switching noise from coupling into analog FSR traces.

## 3.1 Kinematic Vibration Sensor — MPU-6050

The MPU-6050 is a MEMS-based 6-axis IMU (3-axis accelerometer + 3-axis gyroscope) mounted internally near the writing tip to maximise sensitivity to vertical tremor oscillations in the 4–6 Hz Parkinsonian frequency range. It communicates via I2C at up to 400 kHz.

### Electrical Connections

| MPU-6050 Pin | Connects To | Notes |
| --- | --- | --- |
| VDD (pin 13) | 3.3 V rail | Add 100 nF ceramic decoupling cap physically close to this pin |
| GND (pin 18) | GND | System ground |
| SDA (pin 23) | GPIO 8 | 4.7 kΩ pull-up to 3.3 V · verify breakout board does not already include pull-ups |
| SCL (pin 24) | GPIO 9 | 4.7 kΩ pull-up to 3.3 V · verify breakout board does not already include pull-ups |
| AD0 (pin 11) | GND | Sets I2C address to 0x68 · must not be left floating |
| INT, FSYNC | No Connect (NC) | Leave unconnected · pull to GND via 10 kΩ if noise is observed |
| VLOGIC (pin 8) | 3.3 V rail | Tie to 3.3 V on variants that expose this pin |
| CLKIN (pin 2) | NC | Internal oscillator used |
| AUX_DA, AUX_CL | NC | Secondary I2C bus not used in this design |
| RESV pins | NC | Reserved — do not connect |

| i | NOTE: The 100 nF VCC decoupling capacitor on MPU-6050 pin 13 must be placed within 1 mm of the pin pad on a PCB layout. At greater distances it becomes ineffective at suppressing BLE radio-frequency interference on the 3.3 V rail. |
| --- | --- |

## 3.2 Tri-FSR Analog Front-End

Three Force Sensitive Resistors capture pressure dynamics without requiring bulky operational amplifiers or load cells. Each FSR is configured in a passive voltage divider circuit with identical component values, forming three independent ADC1 input channels.

### Sensor Selection Rationale

| Sensor | Active Area | GPIO | Placement | Selection Rationale |
| --- | --- | --- | --- | --- |
| FSR-400 | 7.5 mm diameter | GPIO 1 (ADC1_CH1) | Inside barrel — ink cartridge distal end | Small area fits the narrow tip channel. A micro-plunger transfers only axial (downward) force, mechanically isolating it from lateral grip pressure. |
| FSR-402 A | 44 mm diameter | GPIO 2 (ADC1_CH2) | Lower barrel — primary finger contact | Large area covers the full fingerprint region regardless of individual grip style. Captures both sustained grip rigidity and high-frequency micro-tremor oscillations. |
| FSR-402 B | 44 mm diameter | GPIO 3 (ADC1_CH3) | Lower barrel — secondary finger contact | Eliminates blind spots when grip contact shifts between users. Two sensors working in tandem provide redundancy and capture pill-rolling micro-tremors regardless of finger positioning. |

### Base Voltage Divider Circuit (identical for all three FSRs)

Each FSR terminal 1 connects to 3.3 V. Terminal 2 connects to Node A — a junction point shared with a 10 kΩ pull-down resistor to GND and a 100 nF capacitor in parallel with the pull-down. As applied force increases, FSR resistance drops from >10 MΩ (unloaded) toward ~10 kΩ (firmly pressed), raising the Node A voltage proportionally from near 0 V toward 3.3 V.

### RC Low-Pass Hardware Filter

The 100 nF capacitor in parallel with the 10 kΩ pull-down resistor forms a passive RC low-pass filter:

fc  =  1 / (2π × 10,000 Ω × 100×10⁻⁹ F)  ≈  159 Hz

This filter passes the 4–6 Hz biological tremor signal with negligible attenuation while rejecting 50/60 Hz mains interference and high-frequency ADC switching noise.

# 4. Option 2 Boot-Safe RC Decoupling — All Three FSR Channels

Because GPIO 1, 2, and 3 are ESP32-S3 strapping pins, a direct voltage divider connection risks holding these pins at an incorrect voltage level during boot, causing boot loops or preventing firmware upload. Option 2 adds a 100 kΩ series resistor and a 10 µF electrolytic boot-delay capacitor between Node A (the FSR divider output) and the GPIO pin (Node B).

## 4.1 How Option 2 Works

- During boot (0–~2 ms):: The ESP32-S3 samples strapping pins to determine startup mode. Node B (the GPIO pin) is held near 0 V by the 10 µF capacitor, which has not yet charged. The ESP32 reads the correct strapping value.

- RC time constant:: τ = 100 kΩ × 10 µF = 1 second. Node B rises slowly toward Node A over approximately 1 second after power-on.

- During normal operation (after ~3 s):: Node B has settled and tracks the FSR pressure signal normally. Firmware enforces a 3-second startup delay before ADC sampling begins.

- 100 kΩ series resistor function:: Makes the FSR divider a high-impedance source from the GPIO's perspective during boot. The ESP32-S3 internal strapping logic easily overrides it.

## 4.2 Complete Circuit Per FSR Channel

The following circuit applies identically to all three FSR channels. Only the FSR type and GPIO pin differ.

FSR-400 (Tip Sensor) → GPIO 1 (ADC1_CH1)

3.3V

│

├── Terminal 1 of FSR-400

│

Terminal 2 of FSR-400

│

├─────────────────────────────── NODE A

│                                      │

[10kΩ pull-down]                   [100kΩ series resistor]

│                                      │

[100nF cap]  ← parallel to 10kΩ    NODE B ──────────────► GPIO 1

│                                      │

GND                               [10µF electrolytic cap]

│ (+ leg to Node B, − leg to GND)

GND

FSR-402 Grip A → GPIO 2 (ADC1_CH2)

3.3V

│

├── Terminal 1 of FSR-402 (A)

│

Terminal 2 of FSR-402 (A)

│

├─────────────────────────────── NODE A

│                                      │

[10kΩ pull-down]                   [100kΩ series resistor]

│                                      │

[100nF cap]  ← parallel to 10kΩ    NODE B ──────────────► GPIO 2

│                                      │

GND                               [10µF electrolytic cap]

│ (+ leg to Node B, − leg to GND)

GND

FSR-402 Grip B → GPIO 3 (ADC1_CH3)

3.3V

│

├── Terminal 1 of FSR-402 (B)

│

Terminal 2 of FSR-402 (B)

│

├─────────────────────────────── NODE A

│                                      │

[10kΩ pull-down]                   [100kΩ series resistor]

│                                      │

[100nF cap]  ← parallel to 10kΩ    NODE B ──────────────► GPIO 3

│                                      │

GND                               [10µF electrolytic cap]

│ (+ leg to Node B, − leg to GND)

GND

## 4.3 Boot Timing Diagram

Time:       0 ms          ~2 ms              ~1000 ms         ~3000 ms

│              │                     │                │

3.3V on: ───►│              │                     │                │

│              │                     │                │

Node A:      │──────────────┼─── rises with FSR divider immediately

│              │                     │                │

Node B:      │──────────────┼─────────────────────┼──► slowly rises (RC τ = 1 s)

(GPIO pin)   │              │                     │                │

│    ESP32-S3 reads                  │                │

│    strapping pins                  │                │

│    Node B ≈ 0V ✓                   │                │

│                                    │                │

│                             Node B settled     Firmware starts

│                                                ADC sampling ✓

| ! | POLARITY WARNING: The 10 µF capacitors are electrolytic and polarised. The positive leg (+) must connect to Node B (the GPIO side). The negative leg (−) must connect to GND. Reversed installation will damage the capacitor and may damage the ESP32-S3. |
| --- | --- |

| i | FIRMWARE NOTE: Add a 3-second startup delay in firmware before beginning ADC sampling. This allows Node B to fully settle after power-on and prevents false readings during the RC charging period. |
| --- | --- |

# 5. Power Management Subsystem

The pen operates entirely untethered. Stability on the 3.3 V rail is critical: any voltage fluctuation corrupts analog FSR readings by shifting the baseline voltage of all three voltage dividers simultaneously, introducing false tremor artefacts that cannot be distinguished from real signal.

| Parameter | Value |
| --- | --- |
| Connector | JST-PH 2-pin (positive + negative only) |
| Battery chemistry | Lithium-polymer (LiPo) |
| Nominal voltage | 3.7 V |
| Capacity | 500 mAh |
| Voltage range | 4.2 V (fully charged) → 3.2 V (depleted) |
| Input connection | VBAT / 5V pad on ESP32-S3 SuperMini |
| Regulated output | 3.3 V (flat) via onboard Low-Dropout (LDO) linear regulator |
| Current draw | ~80–120 mA (active BLE + dual-core DSP) |
| Estimated runtime | 4–6 hours continuous — exceeds standard clinical session duration |

| i | JST CONNECTOR NOTE: The schematic uses a 2-pin JST-PH connector (CN1) for the LiPo battery — positive (+) and negative (−) only. A 4-pin JST-PH would imply a balance connector used in multi-cell packs. This design uses a single-cell LiPo and requires only the 2-pin variant. |
| --- | --- |

# 6. Complete Schematic Pinout & Connection Matrix

This table defines every hardware connection in the system. It reflects all corrections from the schematic review: corrected MCU label (ESP32-S3 not C3), AD0 explicitly grounded, VLOGIC tied to 3.3 V, decoupling cap on MPU-6050 VDD, Option 2 RC decoupling on all FSR channels, and 2-pin JST-PH connector.

| Component | Component Pin | ESP32-S3 Pin / Net | Value / Notes |
| --- | --- | --- | --- |
| LiPo battery | (+) positive | VBAT / 5V pad | JST-PH 2-pin connector · raw 3.7–4.2 V input |
| LiPo battery | (−) negative | GND | System ground |
| MPU-6050 | VDD (pin 13) | 3.3 V | 100 nF ceramic decoupling cap to GND · place within 1 mm |
| MPU-6050 | GND (pin 18) | GND | System ground |
| MPU-6050 | SDA (pin 23) | GPIO 8 | I2C data · 4.7 kΩ pull-up to 3.3 V (verify breakout) |
| MPU-6050 | SCL (pin 24) | GPIO 9 | I2C clock · 4.7 kΩ pull-up to 3.3 V (verify breakout) |
| MPU-6050 | AD0 (pin 11) | GND | Sets I2C address 0x68 · must not float |
| MPU-6050 | VLOGIC (pin 8) | 3.3 V | Tie to 3.3 V · do not leave floating |
| MPU-6050 | INT, FSYNC | NC | Optional: 10 kΩ to GND if EMI observed |
| MPU-6050 | RESV, AUX_DA, AUX_CL, CLKIN | NC | Reserved / unused · do not connect |
| FSR-400 (tip) | Terminal 1 | 3.3 V | Active power |
| FSR-400 (tip) | Terminal 2 | Node A | Junction: 10 kΩ to GND ‖ 100 nF to GND |
| R (series) | Node A → Node B | 100 kΩ | Boot-safe series resistor (Option 2) |
| C (boot delay) | Node B → GND | 10 µF electrolytic | Boot delay cap · (+) to Node B · (−) to GND |
| FSR-400 (tip) | Node B | GPIO 1 (ADC1_CH1) | Final GPIO connection after RC network |
| FSR-402 grip A | Terminal 1 | 3.3 V | Active power |
| FSR-402 grip A | Terminal 2 | Node A | Junction: 10 kΩ to GND ‖ 100 nF to GND |
| R (series) | Node A → Node B | 100 kΩ | Boot-safe series resistor (Option 2) |
| C (boot delay) | Node B → GND | 10 µF electrolytic | Boot delay cap · (+) to Node B · (−) to GND |
| FSR-402 grip A | Node B | GPIO 2 (ADC1_CH2) | Final GPIO connection after RC network |
| FSR-402 grip B | Terminal 1 | 3.3 V | Active power |
| FSR-402 grip B | Terminal 2 | Node A | Junction: 10 kΩ to GND ‖ 100 nF to GND |
| R (series) | Node A → Node B | 100 kΩ | Boot-safe series resistor (Option 2) |
| C (boot delay) | Node B → GND | 10 µF electrolytic | Boot delay cap · (+) to Node B · (−) to GND |
| FSR-402 grip B | Node B | GPIO 3 (ADC1_CH3) | Final GPIO connection after RC network |

# 7. PCB Layout Considerations

When transitioning from the breadboard prototype to a custom PCB fitted inside the pen chassis, adhere strictly to the following constraints.

## 7.1 Antenna Keep-Out Zone

The ESP32-S3 SuperMini PCB antenna must not be covered by any copper plane, metallic housing, or conductive shielding. Define a keep-out zone extending at least 3 mm beyond the antenna traces in all directions. No metallic pen barrel components should sit directly above or below the antenna region.

## 7.2 Analog / Digital Trace Isolation

Route I2C digital traces (SDA, SCL on GPIO 8, 9) as far as possible from the three analog ADC traces (GPIO 1, 2, 3). High-speed I2C clock edges switching at 100–400 kHz can capacitively couple into adjacent FSR traces, producing periodic noise artefacts that mimic micro-tremors. Use a ground-fill guard trace between the digital and analog routing zones where board area permits.

## 7.3 Decoupling Capacitor Placement

| Capacitor | Target | Maximum Distance | Note |
| --- | --- | --- | --- |
| 100 nF (×3, RC filter) | FSR pull-down resistor | 2 mm from ADC pin pad | One per FSR channel — Node A |
| 10 µF (×3, boot delay) | GPIO pin (Node B) | 2 mm from GPIO pin pad | Electrolytic — observe polarity |
| 100 nF (MPU-6050 decoupling) | MPU-6050 VDD pin 13 | 1 mm from VDD pin | Critical — BLE RF suppression |

## 7.4 Ground Plane

Use a solid copper ground pour across the entire PCB bottom layer. This provides a low-impedance return path for analog signals, reduces susceptibility to external EMI, and acts as a passive heat sink for the ESP32-S3 during sustained BLE transmission.

## 7.5 FSR Mechanical Interface

- FSR-400 tip: Requires a dedicated micro-plunger channel machined into the pen barrel. The channel must constrain lateral movement so only axial (downward) force from the ink cartridge reaches the sensor active area.

- FSR-402 grip sensors: Must be conformally adhered to the curved barrel surface using a thin compliant adhesive layer (e.g. silicone transfer tape) to ensure uniform contact pressure across the 44 mm active area without creating rigid pressure hot-spots.

| ! | LIFECYCLE NOTE: Formal lifecycle testing of the plunger mechanism under repeated ballpoint writing impact is required before clinical deployment. High-frequency strokes may degrade the plunger channel geometry over time, altering axial force transfer characteristics and introducing drift in the FSR-400 baseline. |
| --- | --- |

# 8. Signal Integrity & Noise Budget

The ESP32-S3 ADC achieves approximately 11-bit effective resolution in practice. The following layered noise mitigation strategy maintains signal quality for 4–6 Hz tremor detection across all three FSR channels.

| Technique | Stage | Effect |
| --- | --- | --- |
| RC hardware LPF (fc ≈ 159 Hz) | Analog front-end (Node A) | Reduces broadband noise floor before ADC sampling |
| 100 kΩ series + 10 µF boot delay | Boot protection (Node B) | Isolates GPIO from divider during strapping pin sampling |
| Software oversampling (4× average) | Firmware | Recovers ~0.5 effective ADC bits |
| DC offset removal | DSP chain | Eliminates slow drift from FSR viscoelastic creep |
| Band-pass filter 3.5–6.5 Hz | DSP chain | Isolates Parkinsonian tremor band in software |
| ADC1-only routing | Hardware architecture | ADC2 shares BLE radio paths — avoided entirely |
| Digital/analog GPIO separation | PCB layout | Prevents I2C clock coupling into FSR ADC traces |

## 8.1 Frequency Domain Budget

Signal of interest:      4–6 Hz      (Parkinsonian tremor band)

RC filter cutoff:        ~159 Hz     (Node A hardware LPF)

High-pass filter (SW):   1 Hz        (DC offset removal)

Band-pass filter (SW):   3.5–6.5 Hz  (DSP tremor isolation)

FFT window:              1024 samples @ 1 kHz = 1.024 s

FFT frequency resolution: ~0.98 Hz/bin

ADC sample rate:         1 kHz per channel

# 9. Complete Component Bill of Materials

| Ref | Component | Value / Part No. | Qty | Purpose |
| --- | --- | --- | --- | --- |
| U1 | ESP32-S3 SuperMini | ESP32-S3 (dual-core LX7) | 1 | MCU · BLE · ADC · LDO |
| U2 | IMU | MPU-6050 | 1 | 6-axis MEMS IMU · I2C |
| U3 | Tip force sensor | FSR-400 | 1 | Axial writing force · 7.5 mm |
| U4, U5 | Grip force sensors | FSR-402 | 2 | Fingerprint grip · 44 mm |
| BT1 | LiPo battery | 3.7 V 500 mAh | 1 | Main power · JST-PH 2-pin |
| CN1 | Battery connector | JST-PH 2-pin | 1 | LiPo interface · 2-pin only |
| H1 | Tip FSR header | 2-pin male header | 1 | FSR-400 connection |
| H4, H5 | Grip FSR headers | 2-pin male header | 2 | FSR-402 A and B connections |
| R1, R5, R8 | FSR pull-down resistors | 10 kΩ | 3 | Voltage divider lower leg (Node A) |
| R3, R7, R9 | Boot-safe series resistors | 100 kΩ | 3 | Option 2 · Node A to Node B isolation |
| R2, R6 | I2C pull-up resistors | 4.7 kΩ | 2 | SDA + SCL pull-ups (if not on breakout) |
| C1, C4, C7 | FSR RC filter capacitors | 100 nF ceramic | 3 | Hardware LPF · parallel to 10 kΩ pull-down |
| C3, C6, C8 | Boot delay capacitors | 10 µF electrolytic | 3 | Option 2 · Node B to GND · observe polarity |
| C2 | MPU-6050 decoupling cap | 100 nF ceramic | 1 | VDD pin 13 · place within 1 mm |
| R4 | Floating — delete | 100 kΩ | 0 | REMOVE — unconnected in schematic v1 |

# 10. Schematic Review — Corrections Applied (v1 → v2)

The following issues were identified in schematic v1 and corrected in this v2 document.

| # | Severity | Issue | Correction Applied |
| --- | --- | --- | --- |
| 1 | Critical | MCU labeled 'ESP32-C3 SuperMini' — wrong chip | Corrected to ESP32-S3 SuperMini throughout |
| 2 | Critical | GPIO 1, 2, 3 are strapping pins — direct FSR divider connection risks boot loops | Option 2 RC decoupling applied: 100 kΩ series + 10 µF boot delay cap per channel |
| 3 | Critical | MPU-6050 VDD decoupling capacitor missing | 100 nF ceramic cap added to VDD pin 13 |
| 4 | Warning | R4 (100 kΩ) floating — not connected | R4 removed from BOM and schematic |
| 5 | Warning | AD0 pin shown as NC — must not float | AD0 explicitly tied to GND · I2C address fixed at 0x68 |
| 6 | Warning | VLOGIC pin disposition not specified | VLOGIC tied to 3.3 V rail |
| 7 | Warning | CN1 shown as 4-pin connector — incorrect for single-cell LiPo | Changed to 2-pin JST-PH (positive + negative only) |
| 8 | Info | I2C net labels (I2C_SDA, I2C_SCL) ambiguity on ESP32 block | Confirmed: I2C_SDA → GPIO 8 · I2C_SCL → GPIO 9 · documented in pinout table |

# 11. Ethical & Data Safety Considerations

- BLE encryption: All BLE transmission must use AES-128 encryption before any clinical deployment to prevent interception of patient kinematic data.

- On-device storage: The companion mobile application must use encrypted on-device storage (SQLCipher on Android · iOS Data Protection class NSFileProtectionComplete).

- Cloud synchronisation: Any cloud sync of feature vectors must comply with HIPAA (US) or GDPR (EU), including explicit informed patient consent and data minimisation.

- Clinical representation: The device must never be represented to patients or clinicians as a diagnostic instrument. Results must always accompany professional clinical assessment.

- Data retention: Feature vectors and raw session logs must be subject to a defined retention policy agreed with the institution's data protection officer before deployment.

# 12. Revision History

| Version | Date | Changes |
| --- | --- | --- |
| v1.0 | June 2026 | Initial hardware architecture document |
| v2.0 | June 2026 | Corrected MCU label (C3→S3) · Added Option 2 RC boot-safe decoupling on all FSR channels · Added MPU-6050 VDD decoupling cap · Grounded AD0 · Tied VLOGIC to 3.3 V · Removed floating R4 · Corrected CN1 to 2-pin JST-PH · Full schematic corrections table added |

For signal processing and firmware documentation see SIGNAL_PROCESSING.md  ·  For mobile application see MOBILE_APP.md